# Slocum Glider Check-IN

# DATE:      ___________

# GLIDER:  ___________	SB: ____________

Vehicle Powered

Power on vehicle in order to fully retract pump, and/or to deflate air bladder.

Vehicle Cleaning (hose down with pressure)

### Nose cone

- Remove nose cone

- Loosen altimeter screws, and remove altimeter or leave temporarily attached

- Retract pump

- Remove altimeter and hose diaphragm removing all sand, sediment, bio oils

- Clean nose cone and altimeter

### Tail cone

- Remove tail cone

- Hose and clean anode and air bladder making sure air bladder is completely clean

- Clean cowling

### Wing rails

- Remove wing rails and hose down

### Tail plug cleaning

- Dip red plug in alcohol and clean plug if especially dirty

- Re-dip red plug and repeatedly insert and remove to clean the glider plug

- Compress air glider female connector

- Lightly silicon red plug and replace in glider once silicon has been dispersed evenly in the plugs

## CTD Comparison Check

- Inspect CTD sensor for any sediment buildup, take pictures of anything suspicious or make note.

Static Tank Test

SBE19						Glider (SBE41CP or pumped unit)

Temperature:	________________		Temperature:	________________

Conductivity:	________________		Conductivity:	________________

### CTD Maintenance if comparison is not acceptable (reference SeaBird Application Note 2D)

- Perform CTD backward/forward flush with 1% Triton X-100 solution

- Perform CTD backward/forward flush with 500 – 1000 ppm bleach solution

- Perform the same on a pumped unit, just different approach

- Repeat comparison test if  above results not within T < .01 C, C < .005 S/m

SB19						Glider (SB41CP or pumped unit)

Temperature:	________________		Temperature:	________________

Conductivity:	________________		Conductivity:	________________

Vehicle Disassembled

- Check leak points for water or salt buildup

- BACKUP FLASH CARDS in /coolgroup/gliderData/glider_OS_backups/<glider>/<glider-deploymentID>/<from glider>,<from sb_0xxx> 
________________________________________________________________
DO NOT DELETE DATA OFF CARDS

- Change permissions on <glider-deploymentID> folder to read, write, execute for owner and group, and read, execute for everyone

- Remove used batteries and place in return crate
________________________________

- Re-assemble glider with a vacuum